Title: Alginate Over Metallic Structures in CBCT Do Not Reduce Artifact

Objective: This study evaluates the impact of alginate impression material on Metal artifacts in CBCT imaging, specifically focusing on image quality and its potential interference with anatomical visualization around dental implants.

Methods: An ex-vivo experiment was conducted using dry human skulls with three Straumann titanium dental implants placed in edentulous sites. Ten CBCT scans (Scans A-J) were acquired with and without the use of alginate impression material. Two board-certified radiologists performed a qualitative assessment of scatter artifacts using a modified Likert scale. A quantitative evaluation was also conducted to analyze pixel intensity values (PIVs) in areas surrounding the implants, focusing on the impact of alginate on scatter artifacts.

Results: Scan A (control) exhibited the least amount of scatter artifacts, serving as the baseline for comparison. Scans B, E, G, and I (without alginate) displayed PIVs and scatter levels comparable to the control, indicating lower distortion. In contrast, scans that used alginate (C, D, F, H, J) demonstrated significantly higher PIVs with differences ranging from -411.31 to -516.08 in mean pixel intensities (p-values between 0.000 and 0.091). The increased scatter in alginate-based scans suggests that the dense, calcium-based composition of alginate might contribute to elevated PIVs around metallic implants.

Qualitative analysis, however, revealed that alginate improved the overall image quality by softening the scans but reduced sharpness around implant areas. This suggests a trade-off between image sharpness and artifact clarity when using alginate

Conclusion: Under the tested conditions, the use of alginate impression material increased PIVs in the vicinity of dental implants in CBCT imaging. These findings imply that rather than reducing scatter, alginate may exacerbate artifacts, particularly in metal-dense regions, thereby reducing image quality. Further research is needed to explore alternative impression materials or techniques to improve image quality and minimize metal artifacts in CBCT imaging.

| Comparison of all scans acquired with Scan A (Control Scan) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Comparison | Mean Pixel Intensity (Scan A vs X) | Mean Difference | p-value | Significance |
| Scan A (Control) | 123.05 vs — | — | — | — |
| Scan A vs Scan C (Alginate) | 123.05 vs 534.36 | -411.31 | 0.091 | Not significant |
| Scan A vs Scan D (Alginate) | 123.05 vs 589.17 | -466.12 | 0.029 | Significant |
| Scan A vs Scan F (Alginate) | 123.05 vs 632.21 | -509.16 | 0.001 | Significant |
| Scan A vs Scan H (Alginate) | 123.05 vs 598.67 | -475.62 | 0.003 | Significant |
| Scan A vs Scan J (Alginate) | 123.05 vs 639.13 | -516.08 | 0.000 | Highly significant |
| Scan A vs Scan B (No Alginate) | 123.05 vs 395.84 | -272.79 | 0.622 | Not significant |
| Scan A vs Scan E (No Alginate) | 123.05 vs 261.09 | -138.04 | 0.979 | Not significant |
| Scan A vs Scan G (No Alginate) | 123.05 vs 465.33 | -342.28 | 0.123 | Not significant |
| Scan A vs Scan I (No Alginate) | 123.05 vs 372.97 | -249.92 | 0.453 | Not significant |

| Comparison of Scans with similar Scenarios – Presence Vs Absence of Alginate |  |  |  |  |
| --- | --- | --- | --- | --- |
| Comparison | Mean Pixel Intensity (Scan X vs Y) | Mean Difference | p-value | Significance |
| Scan E (No Alginate) vs Scan F(Alginate) | 261.09 vs 632.21 | -371.12 | 0.012 | Significant |
| Scan G (No Alginate) vs Scan H(Alginate) | 465.33 vs 598.67 | -133.34 | 0.087 | Not significant |
| Scan I (No Alginate) vs Scan J(Alginate) | 372.97 vs 639.13 | -266.16 | 0.049 | Significant |