Title: Alginate Over Metallic Structures in CBCT Do Not Reduce Artifact

Objective: To characterize and compare peri-implant CBCT gray-value distributions at standardized sampling locations around titanium implants in matched scan configurations acquired with and without alginate, and to assess consistency across three trials and tray status.

Methods: An ex-vivo experiment was conducted using dry human skulls with three Straumann titanium dental implants placed in edentulous sites. Ten CBCT scans (Scans A-J) were acquired with and without the use of alginate impression material. Two board-certified radiologists performed a qualitative assessment of scatter artifacts using a modified Likert scale. A quantitative evaluation was also conducted to analyze pixel intensity values (PIVs) in areas surrounding the implants, focusing on the impact of alginate on scatter artifacts.

Results: The spreadsheet audit confirmed 18 site rows with 72 measurements per site, organized as three trials of 24 measurements. Nine of ten manuscript scan means reproduced from trial 1 at two decimals; Scan F was 632.2047 in the spreadsheet (632.20) versus 632.21 in the manuscripts. Across the matched B–C, E–F, G–H, and I–J configurations, alginate scans had higher descriptive mean peri-implant gray values in each trial. Because repeated points within a scan are correlated, these summaries are exploratory and do not independently establish artifact magnitude.

Qualitative analysis, however, revealed that alginate improved the overall image quality by softening the scans but reduced sharpness around implant areas. This suggests a trade-off between image sharpness and artifact clarity when using alginate

Conclusion: In this ex-vivo dataset, matched scans acquired with alginate showed higher descriptive peri-implant gray values across three trials. CBCT gray values are protocol-dependent surrogate imaging outcomes, and repeated sampling locations within a scan are not independent experimental replicates. These findings should not be interpreted as demonstrating that alginate reduces or worsens metal artifact. Independent acquisition-level replication and a validated direct artifact outcome are needed for confirmatory inference.

| Comparison of all scans acquired with Scan A (Control Scan) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Comparison | Mean Pixel Intensity (Scan A vs X) | Mean Difference | p-value | Significance |
| Scan A (Control) | 123.05 vs — | — | — | — |
| Scan A vs Scan C (Alginate) | 123.05 vs 534.36 | -411.31 | 0.091 | Not significant |
| Scan A vs Scan D (Alginate) | 123.05 vs 589.17 | -466.12 | 0.029 | Significant |
| Scan A vs Scan F (Alginate) | 123.05 vs 632.21 | -509.16 | 0.001 | Significant |
| Scan A vs Scan H (Alginate) | 123.05 vs 598.67 | -475.62 | 0.003 | Significant |
| Scan A vs Scan J (Alginate) | 123.05 vs 639.13 | -516.08 | 0.000 | Highly significant |
| Scan A vs Scan B (No Alginate) | 123.05 vs 395.84 | -272.79 | 0.622 | Not significant |
| Scan A vs Scan E (No Alginate) | 123.05 vs 261.09 | -138.04 | 0.979 | Not significant |
| Scan A vs Scan G (No Alginate) | 123.05 vs 465.33 | -342.28 | 0.123 | Not significant |
| Scan A vs Scan I (No Alginate) | 123.05 vs 372.97 | -249.92 | 0.453 | Not significant |

| Comparison of Scans with similar Scenarios – Presence Vs Absence of Alginate |  |  |  |  |
| --- | --- | --- | --- | --- |
| Comparison | Mean Pixel Intensity (Scan X vs Y) | Mean Difference | p-value | Significance |
| Scan E (No Alginate) vs Scan F(Alginate) | 261.09 vs 632.21 | -371.12 | 0.012 | Significant |
| Scan G (No Alginate) vs Scan H(Alginate) | 465.33 vs 598.67 | -133.34 | 0.087 | Not significant |
| Scan I (No Alginate) vs Scan J(Alginate) | 372.97 vs 639.13 | -266.16 | 0.049 | Significant |